Samenvatting Harris 21.

21.1 – 21.5.

21.1)

- / 463. Hollow cathode lamp (cathode bombarded energetic Ne+/Ar+ → gaseous atoms → electrons → excited “analyte”atom → emission desired λ). Small wavelengths (0.001nm vs 100nm) & sharp bandwidt h atoms.

- Atomic fluorescence (rare) vs absorption, atomic emission (without lamp, collisions plasma → promote → emission).

21.2)

- Premix burner (fuel, oxidant & sample mixed), nebulization (small droplets), aerosol (suspension liquid or solid in gas).

- Rich (↑fuel: excess carbon reduces metal oxides → ↑sensitivity) vs lean (excess oxidant → ↑T) flame. Graphite furnace: hollow cathode measurement; ↑residence time → ↑sensitivity.

- Furnace: drying (20s 125˚C), charring (aka pyrolysis; 60s 400˚C destroying organic matter), atomized (10s 2100˚C), cleaned (3s 2500˚C), purged Ar/N2 each step except atomization. Transversely heating (side to side; uniform). Coated →| Memory effect (previous samples).

- Preconcentration (evaporate multiple samples before procedure). Direct solid sampling (no preparation). Matrix (everything other than analyte), matrix modifier (reduce loss analyte → volatility).

- Inductivity coupled plasma (→ ↑T, stabile, inert Ar environment; fig 21.12).

21.3)

- T → breakdown sample & state. Degeneracies (g* & g0; number of states energy level). Boltzman distribution: N*/N0 = g*/g0∙e-ΔE/kT.

- T↕ → (1) Atomic absorption: insignificant effect, (2) emission: 4% Δ10K (→ stabile flame; usually plasma).

Linewidth emission (#λ!) governed Heisenberg uncertainty principle: lifetime excited state↓ → ↕↕energy [uncertainty↑]: δEδt  h / 4π δE: uncertainty energy difference ground & excited states; δt: lifetime before decay groundstate; h: planck’s constant. δE = “Δ from ΔE”. Relative uncertainty = δE/ΔE.

- δE/ΔE = δλ/λ. Further broadening by Doppler effect (atoms moving; δλ ≈ λ(7∙10-7)√(T/M); M mass in u) and pressure broadening: collisions shorten lifetime.

- (..)

- Beam chopping (blocking lamp) to obtain background, but not scatter, scatter background by deuterium lamp background correction (..) or Zeeman correction system (..).

- Detection limit (amount → twice background).

21.5)

  - Interference (effect → ↕analyte signal): (1) spectral interference (overlap other signals), (2) chemical interference (decreases degree atomization analyte; releasing agents to counter), (3) ionization interference (↑T → atoms get ionized; counter ionization supressor).

  - Inductivity coupled plasma hotter (atomization↑) & residence time↑ → enhanced signal. Flame: less uniform, cooler outside, there self-absorption emission inside, plasma uniform.